## Introduction
Introduction
The aim of this exercise is to undertake a basic system-level failure mode effects analysis (FMEA).
This worksheet contains several tabs:
Introduction | This tab
Hazardous Scenarios | List of possible hazards with Severity and P2 scoring listed.
Risk Estimation Criteria | Summary of the different scoring levels and how they link together
Risk Assessment | The main tab to undertake the analysis
HAZOP checklist | A checklist of prompts to stimulate ideas about how things could fail
Task 1
Draw schematic representation of system/subsystem
List the primary functions that the system/subsystem has to perform (consider interfaces to other systems/users in this)
Task 2
For one of the following functions:
Motion and control
Hold a tool
Receive commands
Mounting to a base
Set up proceedure
Reach stiffness and loading
Complete columns A-D, G, H and N in the Risk Assessment worksheet. That means:
Think about possible failure modes (consider HAZOP checklist)
Agree likely Harm
Estimate P1
Propose possible mitigations
Consider at least 3 failure modes and propose mititgations for these

## Risk Assessment
ID | Item / Function | Potential Failure Mode | Potential Cause(s)/Mechanism(s) of Failure | Potential Effect(s) of Failure | Current Control (Prevention) | Hazardous Situation | Occurrence (P1) | Justification | Severity | Occurance (P2) | Score | Criticality | Recommended Action / mitigation | Notes
Unique line item for ease of reference (don’t change once assigned!) | What function does the system/sub-system need to perform | How could it fail to function correctly? | Why might the system/sub-system have failed in this way? | What is the possible impact if it were to fail? | What features/means are currently inplace to reduce the severity and/or probability of occurance? | What Harm might result if this failed? | How likely is this failure to occur? | On what basis / evidence have you chosen this P1 value? | Auto (Hazardous Situation) | Auto (Hazardous Situation) | Auto (P1xP2xS) | Auto (P1xP2xS) | What can you do to reduce the probability/severity of this situation?
EXAMPLE-01 | Move instrument in response to input | No motion when commanded | No power to motors, inadequate interaction of motor shaft to arms | Complete surgery using alternative means |  | HS-02 Long delay / conversion | Occasional |  | Negligible | Probable | 3 | Low | Check that motor wires and power adapter are connected and that the power supply is turned on.
EXAMPLE-02 |  | Delay before motion is carried out | For some desired points, there can be a delay of <10 seconds before motion begins |  |  | HS-01 Short delay | Frequent |  | Negligible | Frequent | 2 | Low | Check software for potential causes of a delay.
EXAMPLE-02 |  | Unintended / incorrect motion when commanded | Failure of encoder |  |  | HS-03 Small uncontrolled motion at tool tip | Frequent |  | Negligible | Probable | 4 | Low | Check that the encoders are accurate, and if not then replace the motor.
EXAMPLE-03 |  | Unintended motion when not commanded | Software failure |  |  | HS-04 Large uncontrolled motion at tool tip | Occasional |  | Negligible | Probable | 3 | Low | Likely an operation issue or a software issue.
EXAMPLE - 04 | Circuitry / Power supply to robot | Excessive heating of electrical components | Short circuit |  |  | HS-12 Thermal injury to operator | Occasional |  | Minor | Probable | 6 | Low | Wrap the exposed metal at the ends of the wires in tape to prevent short circuits. Check for short circuits using multimeter. | The exposed metal of connecting wires has been wrapped in tape.
EXAMPLE - 05 |  | Electrical shock | Improper use of electrical equipment (e.g. introduction of water into circuit), broken electrical equipment |  |  | HS-10 Operator exposed to electrical voltage / current | Remote |  | Critical | Probable | 8 | Medium | Keep liquids far away from robot. Turn the electrical power off while plugging in adapter. | Food and drink is not allowed in the lab, reducing the chance of occurrence (P1).
EXAMPLE-06 | 3D printed components | Injury to operator | Splinters/sharp remnants of 3D printed support structure |  |  | HS-16 Operator injury | Occasional |  | Minor | Occasional | 4 | Low | Remove support using tools before implementing to robot.
EXAMPLE-07 | Base | Robot topples over | Unstable base, forceful collision of tool with environment |  |  | HS-17 Structural failure of robot | Occasional |  | Minor | Probable | 6 | Low | Consider the risk of interaction with environment before running the robot.
EXAMPLE-08 | Arm 1 | Unintended tilt of tool tip | Unstable integration of arm 1 to base and motor 1 shaft |  |  | HS-18 Unintended orientation of tool tip | Frequent |  | Negligible | Probable | 4 | Low | Redesign arm 1 to include counter-balance, or fix arm 1 to the base with more precision. | Can cause systemic innacuracy at the tool tip.
EXAMPLE-09 | Arm 2 | Unintended tilt of tool tip | Unstable integration of arm 2 to arm 1 and motor 2 shaft, unstable fixing of pen to arm 2 |  |  | HS-18 Unintended orientation of tool tip | Frequent |  | Negligible | Probable | 4 | Low | Redesign mechanism by which pen is fixed to arm 2. | Can cause systemic innacuracy at the tool tip.
Note on C1: queena.lin: The manner in which the item or operation potentially fails to meet or deliver the intended function or associated requirement
Note on D1: queena.lin: The design deficiency that results in the failure mode. The specific reason for the failure. (found by asking "why" until the root cause is determined.

## Risk Estimation Criteria
Risk Estimation & Evaluation Criteria
Table 1: Severity Estimation Levels (Qualitative) | [6 empty] | Table 2: Probability Estimation Levels (Semi-Quantitative)
Severity Ranking | Score | Severity Description for injuries or damage to the health of people, or damage to property of the environment |  |  |  |  | Category | Score | Description
Catastrophic | 5 | Results in death |  |  |  |  | Frequent | 5 | Occurs twice per day; "Very high probability - Very likely to happen, often, frequent"
Critical | 4 | Results in permanent impairment or life-threatening injury |  |  |  |  | Probable | 4 | Occurs twice per week; "High probability - likely to happen, often, frequent"
Serious | 3 | Results in injury or impairment requiring professional medical intervention |  |  |  |  | Occasional | 3 | Occurs monthly; "Medium probability: can happen, but not frequently"
Minor | 2 | Results in temporary injury or impairment not requiring professional medical intervention (including significant IT security failure) |  |  |  |  | Remote | 2 | Occurs yearly; "Low probability: unlikely to happen, rare, remote"
Negligible | 1 | Inconvenience or temporary discomfort (including minor IT security failure or similar) |  |  |  |  | Improbable | 1 | Occurs within 10+ years; "Very low probability: very unlikely to happen, very rare, remote"
Table 3: Probability of Occurrence of Harm (POH) Estimation
Probability of the Hazard leading to the Hazardous Situation (P1)
|  | Improbable | Remote | Occasional | Probable | Frequent
Probability of a Hazardous Situation Leading to Harm (P2) | Improbable | Improbable | Improbable | Improbable | Improbable | Improbable
| Remote | Improbable | Improbable | Remote | Remote | Remote
| Occasional | Improbable | Remote | Remote | Occasional | Occasional
| Probable | Improbable | Remote | Occasional | Probable | Probable
| Frequent | Improbable | Remote | Occasional | Probable | Frequent
Table 4: Criteria for Risk Acceptability
Severity
|  | Negligible | Minor | Serious | Critical | Catastrophic
Probability of Occurrence of Harm | Frequent | Low | Medium | High | High | High
| Probable | Low | Medium | Medium | High | High
| Occasional | Low | Low | Medium | Medium | High
| Remote | Low | Low | Low | Medium | High
| Improbable | Low | Low | Low | Low | Medium

## HAZOP checklist
HAZOP checklist
A list of key words to help assess hazards and prompt ideas for how things could fail
Guide Word | Meaning
NO OR NOT | Complete negation of the design intent
MORE | Quantitative increase
LESS | Quantitative decrease
AS WELL AS | Qualitative modification/increase
PART OF | Qualitative modification/decrease
REVERSE | Logical opposite of the design intent
OTHER THAN / INSTEAD | Complete substitution
EARLY | Relative to the clock time
LATE | Relative to the clock time
BEFORE | Relating to order or sequence
AFTER | Relating to order or sequence

## Hazardous Scenarios
Hazardous Scenarios
Note these are prefixed with an ID number to aid uniqueness
Hazardous Scenario | Description | Severity | Probability (P2)
HS-01 Short delay | Delay of <2mins while system is rectified | Negligible | Frequent
HS-02 Long delay / conversion | Delay of >2mins while system is rectified; or, system unavailable, operator has to complete proceedure with alternative means | Negligible | Probable
HS-03 Small uncontrolled motion at tool tip | Small motion of tool tip (<10mm) not commanded by the user, user not in control | Negligible | Probable
HS-04 Large uncontrolled motion at tool tip | Large motion of tool tip (>10mm) not commanded by the user, user not in control | Negligible | Probable
HS-05 Small unintended motion at tool tip | Small motion of tool tip (<10mm) not in the direction intended by the user, user in control | Negligible | Probable
HS-06 Large unintended motion at tool tip | Large motion of tool tip (>10mm) not in the direction intended by the user, user in control | Negligible | Probable
HS-07 Excessive force | Excessive force applied to patient leading to injury | Serious | Occasional
HS-08 Incorrect feedback to operator | Operator is given incorrect information leading to errors in surgery | Minor | Probable
HS-09 Patient exposed to electrical voltage / current | Patient exposed to electrical energy leading to electric shock | Critical | Probable
HS-10 Operator exposed to electrical voltage / current | Operator exposed to electrical energy leading to electric shock | Critical | Probable
HS-11 Thermal injury to patient | Patient exposed to high temperatures leading to burn internal/external | Serious | Probable
HS-12 Thermal injury to operator | Operator exposed to high temperatures leading to burn | Minor | Probable
HS-13 Interference with other OR equipment | System causes interference with other systems in the OR, mis-managed response to patient status from OR team | Critical | Remote
HS-14 Break in sterility | Failure to maintain sterile field, patient exposed to infection risk | Serious | Occasional
HS-15 Lose parts inside patient | Tool or system failure leads to parts inside patient | Serious | Occasional
HS-16 Operator injury | Operator injury due to excessive motion of system outside patient; sharp edges | Minor | Occasional
HS-17 Structural failure of robot | Structural collapse or breakage of robot or robot components | Minor | Probable
HS-18 Unintended orientation of tool tip | Tool tip tilt leads to inaccuracy of output | Negligible | Probable